| [COMPANY LOGO] | Shivam Gupta's Pharmaceutical Company ADDRESS LINE, CITY - PIN CODE (COUNTRY) |
| --- | --- |

ANALYTICAL METHOD VALIDATION REPORT

| NAME OF PRODUCT | abcd |
| --- | --- |
| LABEL | EACH TAB CONTAINS: creatin ... 25 mg |
| REPORT NO. 1234 | PAGE 1 OF 12 |

ANALYTICAL METHOD VALIDATION ASSAY FOR

ABCD

| [COMPANY LOGO] | Shivam Gupta's Pharmaceutical Company ADDRESS LINE, CITY - PIN CODE (COUNTRY) |
| --- | --- |

ANALYTICAL METHOD VALIDATION REPORT

| NAME OF PRODUCT | abcd |
| --- | --- |
| LABEL | EACH TAB CONTAINS: creatin ... 25 mg |
| REPORT NO. 1234 | PAGE 2 OF 12 |

# CONTENTS

| CONTENT | PAGE |
| --- | --- |
| Active Ingredient: creatin | 3 |
| Product and Code | 4 |
| List of Equipment, Materials, Reagents and Reference Standards | 4 |
| Equipment and Instruments | 4 |
| Glass or Other Materials | 5 |
| Reagents | 6 |
| Reference Products | 7 |
| Results of Each Validation Parameter Evaluated, Statistical Calculations and Acceptance Criteria | 8 |
| Discussion of the Results | 9 |
| Suitability of the System | 9 |
| Selectivity of the Method: Interference and Stress | 9 |
| Precision | 10 |
| Linearity of the System and The Method | 10 |
| Accuracy of the Method | 11 |
| Robustness | 11 |
| Range | 11 |
| Conclusions | 12 |
| Post-Approval | 12 |

| [COMPANY LOGO] | Shivam Gupta's Pharmaceutical Company ADDRESS LINE, CITY - PIN CODE (COUNTRY) |
| --- | --- |

ANALYTICAL METHOD VALIDATION REPORT

| NAME OF PRODUCT | abcd |
| --- | --- |
| LABEL | EACH TAB CONTAINS: creatin ... 25 mg |
| REPORT NO. 1234 | PAGE 3 OF 12 |

# Active Ingredient: creatin

Chemical Structure and/or Molecular Weight:

| Molecular Weight | Molecular Formula |
| --- | --- |
| 131.13 | C4H9N3O2 |

| [COMPANY LOGO] | Shivam Gupta's Pharmaceutical Company ADDRESS LINE, CITY - PIN CODE (COUNTRY) |
| --- | --- |

ANALYTICAL METHOD VALIDATION REPORT

| NAME OF PRODUCT | abcd |
| --- | --- |
| LABEL | EACH TAB CONTAINS: creatin ... 25 mg |
| REPORT NO. 1234 | PAGE 4 OF 12 |

# Product and Code

The procedure was carried out as indicated in the 1234 protocol, for the tests: Identification by TITRATION, Assessment by TITRATION for the active ingredient creatin in the product abcd.

# List of Equipment, Materials, Reagents and Reference Standards

## Equipment and Instruments:

The qualification or maintenance reports of the equipment and/or instruments mentioned, as appropriate, are mentioned below:

| Equipment Name | glass |
| --- | --- |
| Identification Code | 123 |
| Brand | 1234 |
| Verification, Calibration and/or Maintenance | 12 |
| Calibration, Verification and/or Maintenance date: | 2025-11-07 00:00:00.000000 |
| Next Calibration, Verification and/or Maintenance date: | 2025-11-11 00:00:00.000000 |

| [COMPANY LOGO] | Shivam Gupta's Pharmaceutical Company ADDRESS LINE, CITY - PIN CODE (COUNTRY) |
| --- | --- |

ANALYTICAL METHOD VALIDATION REPORT

| NAME OF PRODUCT | abcd |
| --- | --- |
| LABEL | EACH TAB CONTAINS: creatin ... 25 mg |
| REPORT NO. 1234 | PAGE 5 OF 12 |

## Glass or Other Materials

The glass or other materials used are detailed below:

| Glass Materials | Characteristics |
| --- | --- |
| glass | 10 ml |
| Other Materials |  |

| [COMPANY LOGO] | Shivam Gupta's Pharmaceutical Company ADDRESS LINE, CITY - PIN CODE (COUNTRY) |
| --- | --- |

ANALYTICAL METHOD VALIDATION REPORT

| NAME OF PRODUCT | abcd |
| --- | --- |
| LABEL | EACH TAB CONTAINS: creatin ... 25 mg |
| REPORT NO. 1234 | PAGE 6 OF 12 |

## Reagents:

The reagents used are detailed below:

| Reagent Name/Brand | Batch | Expiration Date |
| --- | --- | --- |
| hplc | 4566-45 | 2025-11-06 |

| [COMPANY LOGO] | Shivam Gupta's Pharmaceutical Company ADDRESS LINE, CITY - PIN CODE (COUNTRY) |
| --- | --- |

ANALYTICAL METHOD VALIDATION REPORT

| NAME OF PRODUCT | abcd |
| --- | --- |
| LABEL | EACH TAB CONTAINS: creatin ... 25 mg |
| REPORT NO. 1234 | PAGE 7 OF 12 |

## Reference Products:

Standard Data

| Standard Type | Secondary |
| --- | --- |
| Standard Name | 123 |
| Code | 12398 |
| Potency | 78 |
| Due Date of Standardization | 2025-11-11 |

| [COMPANY LOGO] | Shivam Gupta's Pharmaceutical Company ADDRESS LINE, CITY - PIN CODE (COUNTRY) |
| --- | --- |

ANALYTICAL METHOD VALIDATION REPORT

| NAME OF PRODUCT | abcd |
| --- | --- |
| LABEL | EACH TAB CONTAINS: creatin ... 25 mg |
| REPORT NO. 1234 | PAGE 8 OF 12 |

# Results of Each Validation Parameter Evaluated, Statistical Calculations and Acceptance Criteria

The final results obtained in the identification and Assessment tests, in relation to the acceptance criteria are shown below in tables.

Product: abcd
Active Ingredient: creatin
Methodology Used: TITRATION

Method Conditions:

• Detection Wavelength: 254

## Table: System Suitability Test - TITRATION

| Parameters | Variables | Acceptance Criteria | Results |
| --- | --- | --- | --- |
| System Suitability | Retention Time (RT) CV | RT < 2.00% | 0.08% |
|  | Area CV | CV < 2.00% | 0.2% |
|  | Tailing Factor (T) | 0.80 < T < 2.00 | 1.17 |

## Specificity/Interference

| Interference | Acceptance Criteria | Results |
| --- | --- | --- |
| Mobile Phase | ≤ 0.5% of analyte | No Interference |
| Placebo | ≤ 0.5% of analyte | No Interference |
| Placebo + Active | ≤ 0.5% of analyte | No Interference |

## Precision Parameter

| System Precision | Acceptance Criteria | Results |
| --- | --- | --- |
| Average Area (X) | -- | 1366806 |
| Coefficient of Variation of Areas | CV < 2.00% | 0.26% |
| Average Retention Time | -- | 4.546 |
| Coefficient of Variation of RT | CV < 1.00% | 0.05% |

| Method Precision | Acceptance Criteria | Results |
| --- | --- | --- |
| Average Concentration (X) | -- | 25.01mg/Tab |
| Standard Deviation (s) | -- | 0.028 |
| Variation Coefficient (%CV) | CV < 2.00% | 0.11% |

## Intermediate Precision

| Parameter | Acceptance Criteria | Results |
| --- | --- | --- |
| Day 1 - Analyst 1 (Average Concentration) | CV < 2.00% | 24.98mg/Tab CV: 0.11% |
| Day 1 - Analyst 2 (Average Concentration) | CV < 2.00% | 25.02mg/Tab CV: 0.09% |
| Global Coefficient of Variation (Day 1) | CV < 3.00% | 0.1% |
| Day 2 - Analyst 1 (Average Concentration) | CV < 2.00% | 24.99mg/Tab CV: 0.12% |
| Day 2 - Analyst 2 (Average Concentration) | CV < 2.00% | 24.98mg/Tab CV: 0.1% |
| Global Coefficient of Variation (Day 2) | CV < 3.00% | 0.11% |
| Global Coefficient of Variation (Day 1 and 2) | CV < 3.00% | 0.12% |

## Linearity Parameter

| System Linearity | Acceptance Criteria | Results |
| --- | --- | --- |
| Slope | -- | 4913310.239 |
| Intercept | -- | -3147426.944 |
| Correlation Coefficient (r) | r > 0.9970 | 0.9999 |
| Determination Coefficient (R²) | R² > 0.9950 | 0.9997 |

| Method Linearity | Acceptance Criteria | Results |
| --- | --- | --- |
| Slope | -- | 4811338.168 |
| Intercept | -- | 1765097.671 |
| Correlation Coefficient (r) | r > 0.9970 | 0.9996 |
| Determination Coefficient (R²) | R² > 0.9950 | 0.9993 |

## Accuracy Parameter

| Level | Acceptance Criteria | Recovery % |
| --- | --- | --- |
| 80% Level | 98.00 - 102.00% | 99.8% |
| 100% Level | 98.00 - 102.00% | 99.1% |
| 120% Level | 98.00 - 102.00% | 100.9% |
| Overall | CV < 2.00% | 99.9% (CV: 0.20%) |

## Robustness Parameter

| Variable | Acceptance Criteria | CV % |
| --- | --- | --- |
| Flow Rate (±10%) | CV < 2.00% | 0.31% / 0.05% |
| Wavelength (±2nm) | CV < 2.00% | 0.50% / 0.41% |
| Column Make | CV < 2.00% | 0.24% / 0.52% |
| Column Temp (±2°C) | CV < 2.00% | 0.13% / 0.80% |

## Range Parameter

| Parameter | Range | Results |
| --- | --- | --- |
| Precision | 100% | Complies |
| System Linearity | 80%-120% | Complies |
| Method Linearity | 80%-120% | Complies |
| Accuracy | 80%-120% | Complies |

## LOD and LOQ Parameter

| Parameter | Acceptance Criteria | Results |
| --- | --- | --- |
| Limit of Detection (LOD) | Signal to Noise ratio ≥ 3:1 | 0.045 μg/ml |
| Limit of Quantification (LOQ) | Signal to Noise ratio ≥ 10:1 | 0.15 μg/ml |
| Method Used | Signal to Noise Method | Compliant |

## LOD and LOQ Precision Parameter

| Parameter | Acceptance Criteria | Results |
| --- | --- | --- |
| LOD Precision (n=6) | RSD ≤ 20% | 17.3% |
| LOQ Precision (n=6) | RSD ≤ 15% | 9.0% |
| Overall Assessment | Within Acceptance Criteria | Compliant |

| [COMPANY LOGO] | Shivam Gupta's Pharmaceutical Company ADDRESS LINE, CITY - PIN CODE (COUNTRY) |
| --- | --- |

ANALYTICAL METHOD VALIDATION REPORT

| NAME OF PRODUCT | abcd |
| --- | --- |
| LABEL | EACH TAB CONTAINS: creatin ... 25 mg |
| REPORT NO. 1234 | PAGE 9 OF 12 |

# Discussion of the Results

According to the parameters in each of the test: Identification and Assessment of the Active ingredient creatin, the results of the following parameters are discussed below:

## Suitability of the System:

By running the standard during the entire test, we verified the parameters evaluated, in the TITRATION method, for the assessment of the Active ingredient creatin.

• Retention Time (RT) < 2.0%
• Coefficient of Variation %CV < 2.00%
• Asymmetry(T) 0.80 < T < 2.00

Therefore the system suitability test is defined for the corresponding routine test.

## Selectivity of the Method: Interference and Stress

The selectivity parameter (interference of the method) was evaluated for the identification test and assessment of the Active ingredient creatin, applying the TITRATION technique. It was verified that when injecting the possible interferents: the Mobile Phase, Placebo and Placebo + creatin in the equipment, no responses (areas) were obtained in the placebo at the same retention time of the analyte that was obtained in the standard solution and sample solution, so it does not cause any interference in the assessment of the active ingredient under study. Therefore the acceptance parameters are met, so it can be concluded that there is no interference.

Likewise, for the validation trial, the selectivity parameter was evaluated: Standards stress, placebo stress and finished pharmaceutical product stress according to the conditions were evaluated against the readings emitted (areas) under normal conditions of the active ingredient under study. Furthermore, in the chromatograms it was verified that there is no interference with the degradation products, therefore the method is selective.

| [COMPANY LOGO] | Shivam Gupta's Pharmaceutical Company ADDRESS LINE, CITY - PIN CODE (COUNTRY) |
| --- | --- |

ANALYTICAL METHOD VALIDATION REPORT

| NAME OF PRODUCT | abcd |
| --- | --- |
| LABEL | EACH TAB CONTAINS: creatin ... 25 mg |
| REPORT NO. 1234 | PAGE 10 OF 12 |

## Precision:

When studying the precision of the system in the TITRATION test using the reference standard solution, results were obtained with a relative standard deviation of less than 2.00%, which indicates that the system is repeatable in the analysis.

In the repeatability analysis of the method in the Assessment test, the relative standard deviation of less than 2.00%, this demonstrates the precision of the analytical method.

Likewise, in the Intermediate precision analysis, the results obtained by both analysts on different days gave a coefficient of variation of less than 3.0% for the Assessment test, which demonstrates the reliability of the analytical method.

## Linearity of the System and The Method:

The graph of the response (areas) vs. analyte concentration, demonstrates that the technique by TITRATION in the assay produced a linear response for the analyte.

When applying the linearity test, the correlation coefficient value is greater than 0.9970 for the system Linearity and correlation coefficient value is greater than 0.9970 for the Method linearity hence the analytical method is validated for the linearity parameter.

## Accuracy of the Method:

The accuracy parameters were studied by the addition method of the active ingredient creatin over three different concentrations (80%, 100% and 120%).

For the assay of Assessment of the active ingredient creatin by TITRATION, the percentage of recovery was obtained when evaluating the analyte at three different concentration levels, the overall percentage recovery and coefficient of variation is within the specified limit at all the three levels confirming the method is accurate.

| [COMPANY LOGO] | Shivam Gupta's Pharmaceutical Company ADDRESS LINE, CITY - PIN CODE (COUNTRY) |
| --- | --- |

ANALYTICAL METHOD VALIDATION REPORT

| NAME OF PRODUCT | abcd |
| --- | --- |
| LABEL | EACH TAB CONTAINS: creatin ... 25 mg |
| REPORT NO. 1234 | PAGE 11 OF 12 |

## Robustness:

According to the results obtained in the TITRATION technique and by evaluating the coefficient of variation between the nominal factor and the variable factor, it is shown that there is no significant variation in the concentration of the product.

## Range:

After evaluating the performance characteristics throughout the validation, especially the Linearity, Precision and the Accuracy studies, for the active ingredient creatin, we can conclude that the analytical methodology is validated within the operating range of 80% to 120% for the assessment assay.

| [COMPANY LOGO] | Shivam Gupta's Pharmaceutical Company ADDRESS LINE, CITY - PIN CODE (COUNTRY) |
| --- | --- |

ANALYTICAL METHOD VALIDATION REPORT

| NAME OF PRODUCT | abcd |
| --- | --- |
| LABEL | EACH TAB CONTAINS: creatin ... 25 mg |
| REPORT NO. 1234 | PAGE 12 OF 12 |

# Conclusions

All the Results comply with the acceptance criteria. Hence the Analytical method validation for Assay of abcd by TITRATION has been successfully validated for System Suitability, Specificity, System Precision, Method Precision, Intermediate Precision, Linearity, Recovery, Robustness, Range, Lod Loq, Lod Loq Precision parameters.

Likewise, when evaluating the stability of the sample in the active ingredient creatin, we can indicate there is no significant change in the concentration of the analyte, so the product is considered stable.

# Post-Approval:

This is a specific Report for analytical method validation for Assay of abcd by TITRATION.

This Report has been approved by the following:

|  | Name | Department | Signature | Date |
| --- | --- | --- | --- | --- |
| Prepared By | shivam | Analyst Q.C | [Signature] | 02/11/2025 |
| Checked By | deepak | Asst. Manager Q.C | [Signature] | 02/11/2025 |
| Approved By | saksham | Manager Q.C | [Signature] | 02/11/2025 |